## Slide 1: LOADING EFFECTS IN MEASURING INSTRUMENTS

## Slide 2: What is LOADING EFFECT?
The introduction of the sensing element into the process or system being measured causes the value of the measured variable to change.
The output of a sensor device may deviate from the correct value due to loading effect.

## Slide 3: What is LOADING EFFECT?

## Slide 4: What is LOADING EFFECT?
A
B
A
B
Qt
Qm
Qt  PAB (Hagen-Poisseuille eq)
Depends on Pfrictional
Qm  PAB
Depends on Pfrictional + Porificemeter

## Slide 5: LOADING EFFECT
So introduction of the sensing element into the process or system is always associated with extraction of some energy from the medium which causes the value of the measured variable to change.
This effect cannot be removed altogether – but it can be minimized.

## Slide 6: LOADING EFFECT
MEASUREMENT SYSTEM/ COMPONENT
Input (qi)
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
Primary Input (qi1) | Secondary Input (qi2)
Pressure (P) | Flow (Q)
Force (F) | Velocity (v)
Voltage (E) | Current (i)
Flow (Q) | Pressure (P)
Velocity (v) | Force (F)
Current (i) | Voltage (E)
Power Drain = qi1 qi2
Power Drain
PQ
Fv
Ei
PQ
Fv
Ei

## Slide 7: LOADING EFFECT
Another classification of variables
Flow/Through Variable
Effort/Across Variable
FLOW VARIABLE: an extensive variable in the sense that its magnitude depends on the extent to which the system takes part in energy exchange. Examples: flow, current, velocity
EFFORT VARIABLE: an intensive variable in the sense that its magnitude is independent on the extent to which the system takes part in energy exchange. Examples: pressure, voltage, force
Power Drain is product of a flow variable and an effort variable

## Slide 8: Generalized Input Impedance
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
If qi1 is an effort variable then,
Generalized Input Impedance is defined as
Then power drain/loading effect is
So HIGHER is input impedance Zgi,
LOWER is power drain 

## Slide 9: Example of Generalized Input Impedance
im
Eb
R1
R2
R3
a
b
+

i
Eo
Em
Here qi1 = Em is an effort variable & qi2 = im is a flow variable.
So generalized input impedance is defined as
So now we need to check that to have low loading effect, Rm should be high.
+

Rm
a'
b'
Voltmeter

## Slide 10: Example of Generalized Input Impedance
Rm
a'
b'
im
Eo
Req
a
b
+

Em
By Thevenin’s Theorem
where, q1m  measured value of effort variable (Em)
q1u  true/undisturbed value of effort variable (Eo)
Zgi  input impedance (Rm)
Zgo  output impedance (Req)

## Slide 11: Example of Generalized Input Impedance
Rm
a'
b'
im
Eo
Req
a
b
+

Em
By Thevenin’s Theorem
So, q1m  q1u if (Zgo / Zgi )  0
Or, Zgi >> Zgo
i.e. Rm >> Req
So to have low loading effect, input impedance Rm should be high.

## Slide 12: Generalized Input Admittance
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
If qi1 is a flow variable then,
Generalized Input Admittance is defined as
Then power drain/loading effect is
So HIGHER is input admittance Ygi,
LOWER is power drain 

## Slide 13: Example of Generalized Input Admittance
Here qi1 = im is a flow variable & qi2 = Em is an effort variable.
So now we need to check that to have low loading effect, Ygi should be high.
iu
Eb
R1
R2
R3
a
b
Ammeter
im
Rm
Eab
Req
a
b
Ammeter
im
Rm
By Thevenin’s Theorem
So generalized input and output admittance are defined as

## Slide 14: Example of Generalized Input Admittance
Eab
Req
a
b
Ammeter
im
Rm
So, q1m  q1u if (Ygo / Ygi )  0
Or, Ygi >> Ygo
i.e. Rm << Req
So to have low loading effect, input admittance (1/Rm) should be high.

## Slide 15: Generalized Input Static Stiffness
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
If qi1 is an effort variable and say is force F.
Then generalized Input Impedance is
But then when a steady state is reached while taking reading
v = 0. So Zgi = F/v = F/0 = 
If qi2 is a flow variable and is velocity v.

## Slide 16: Generalized Input Static Stiffness
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
To overcome this difficulty a new parameter is defined as:
So if qi1 = F (effort variable) and qi2 = v (flow variable), then

## Slide 17: Generalized Input Static Stiffness
Generalized input static stiffness is defined as:
where Sgo is the generalized output static stiffness
It can be shown that:
So, q1m  q1u if (Sgo / Sgi )  0
Or, Sgi >> Sgo
So to have low loading effect, input static stiffness Sgi should be high.

## Slide 18: Example of Generalized Input Static Stiffness
A
ks
P
x

P

## Slide 19: Example of Generalized Input Static Stiffness
A
ks
P
x

qi1 = Pm (effort variable)
qi2 = Qm (flow variable)
So to have low loading effect, input static stiffness Sgi should be high i.e. spring stiffness ks should be high.

## Slide 20: Generalized Input Static Compliance
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
If qi1 is a flow variable and say is displacement x.
Then generalized Input Admittance is
If qi2 is an effort variable and is rate of change of force .
But then when a steady state is reached while taking reading

## Slide 21: Generalized Input Static Compliance
MEASUREMENT SYSTEM/ COMPONENT
Output (qo)
Primary Input (qi1)
Secondary Input (qi2)
To overcome this difficulty a new parameter is defined as:
So if qi1 = x (flowvariable) and (effort variable), then
Inverse Stiffness

## Slide 22: Generalized Input Static Compliance
Generalized input static compliance is defined as:
where Cgo is the generalized output static compliance
It can be shown that:
So, q1m  q1u if (Cgo / Cgi )  0
Or, Cgi >> Cgo
So to have low loading effect, input static compliance Cgi should be high.

## Slide 23: Example of Generalized Input Static Compliance
L2
F
Cgo
Gauge Length
L1
Cgi
L1 = CgiF = Lt
L2 = CgoF
Lm = L1 + L2
= (Cgi + Cgo)F
So to have low loading effect, input static compliance Cgi should be high.