## Slide 1: STATIC CHARACTERISTICS OF MEASURING INSTRUMENTS

## Slide 2: Important to know about performance characteristics
Helps in understanding the performance (or rather the lack of performance) of an instrument
Based on this knowledge one gets equipped to select from a set of instruments, the one most suitable for one’s area of application

## Slide 3: What are performance characteristics?
Set of parameters involved in defining the characteristics of the measurement device.
The values against different relevant characteristics of the instrument or measuring system  SPECIFICATIONS
The quality or assurance of required performance of a measuring system is described by the specifications given.
For example – by knowing the accuracy and resolution requirements for your application, you can compute the total error in the measurement device you are considering and verify that it satisfies your needs.

## Slide 4
Performance
Characteristics
STATIC Characteristics
Measurement of quantities are either constant or vary very slowly with time.
Measurements are taken only at steady states.
State between two successive steady states (transition) has no relevance in measurement.
Input-output relationship of the instrument can be described by algebraic equation.
Performance criteria based upon static relations constitute the Static Characteristics.
Thus the static characteristics of a measurement system are those which must be considered when the system or instrument is used under a condition not varying with time.
DYNAMIC Characteristics
Measurements are concerned with rapidly varying quantities.
One examines the dynamic relations which exist between the output and the input.
Done with the help of differential equations or other methods.
Performance criteria based upon dynamic relations constitute the Dynamic Characteristics.

## Slide 5: Example of Static & Dynamic Characteristics
k
c
m
f
x
x = 0
x1
x2
x3
x4
x5
f1
f2
f3
f4
f5
f
x
Static Characteristic
Dynamic Characteristic
x
F = 1; m = 10; c = 2; k = 10

## Slide 6: STATIC CALIBRATION
A process for obtaining the static performance characteristics.
The calibration procedures involve a comparison of the particular characteristic with either a primary standard or a secondary standard – both with a higher accuracy than the instrument to be calibrated, or an instrument of known accuracy.
It checks the instrument against a known standard and subsequently to errors in accuracy.
All measuring instruments must be calibrated against some reference instruments which have a higher accuracy. Thus reference instruments in turn must be calibrated against instrument of still higher grade of accuracy, or against primary standard, or against other standards of known accuracy.
It is essential that any measurement made must ultimately be traceable to the relevant primary standards.
How to obtain the static characteristics?

## Slide 7: STATIC CHARACTERISTICS
Accuracy
Sensitivity
Reproducibility
Drift
Static error
Dead zone
Desirable
Undesirable
Some more quantities have to be defined here which are essential in understanding the above characteristics.

## Slide 8: Scale range and scale span
The choice of proper range of instruments is important in measurement.
In an analogue indicating instrument, the measured value of a variable is indicated on a scale by a pointer.
The region between the limits with in which an instrument is designed to operate for measuring, indicating or recording a physical quantity is called the range of the instruments.

## Slide 9: Scale range and scale span
The Scale Range of an instrument is thus defined as the difference between the largest and the smallest reading of the instrument. Supposing the highest point of calibration is Xmax units while the lowest is Xmin units and the calibration is continuous between the two points, then the instrument range is between Xmin and Xmax .

## Slide 10: Scale range and scale span
Many times it is also said that the instrument range is Xmax. The instrument span is the difference between highest and the lowest point of calibration.
Thus, Span = Xmax - Xmin
For example for a thermometer calibrated between 1000C to 4000C, the range is 1000C to 4000C (or 4000C) but the span is 4000C – 1000C = 3000C.

## Slide 11: Scale range and scale span
For digital instruments, there is another factor that must be considered while determining the range of the instrument. This is the Frequency Range, which is defined as frequencies over which measurements can be performed with a specified degree of accuracy.
For example a moving iron instrument may have a 0-250V range and 0-135Hz frequency range.

## Slide 12: TRUE VALUE
The true value of variable quantity being measured may be defined as the average of an infinite number of measured values when the average deviation due to the various contributing factors tends to zero.
Such an ideal situation is impossible to realize the practice and hence it is not possible to determine the true value of a quantity by experimental means.
The reason for this is that there are several factors such as lags, loading effects, wear or noise pick-up etc.
Normally an experimenter would never know that the value or quantity being measured by experimental means is the true value of the quantity or not.
So one has to remain happy with “an ESTIMATE of the TRUE VALUE”

## Slide 13: ACCURACY
Accuracy is the closeness with which an instrument reading approaches the true value of the quantity being measured.
Thus accuracy of a measurement means conformity to truth.
The accuracy may be expressed as:
  Point accuracy is stated for one or more points in the range, for example, the scale of length may be read within 0.2 mm.
  the instrument is accurate to within x percentage of instrument span at all points on the scale.
  the instrument is accurate to within x percentage of instrument range at all points on the scale.

## Slide 14: PRECISION
Accuracy is many a time confused with Precision.
The term Precise means clearly or sharply defined – implying less variation among multiple readings.
For example an ammeter may possesses high degree of precision
  By virtue of its clearly legible, finely divided, distinct scale and a knife edge pointer with mirror arrangements to remove parallax.
  Due to these the variation in reading will reduce.
  Even though the measurement may be less accurate but will repeat.

## Slide 15: PRECISION
As an example of the difference in meaning of the two terms, suppose above ammeter can read up to 1/100 of an ampere.
Now if its zero adjustment is wrong, every time we take a reading, the readings taken with this ammeter are not accurate, since they do not confirm to truth on account of its faulty zero adjustment.
Though the ammeter is as precise as ever and readings are consistent and clearly defined and can be down to 1/100 of an ampere.
The instrument can be calibrated to remove the zero error.
Thus the accuracy of the instrument can be improved upon by calibration but not the precision.

## Slide 16: PRECISION
In measurement of a set, accuracy is closeness of the measurements to a specific value, while precision is the closeness of the measurements to each other.

## Slide 17: STATIC ERROR
Measurements done with an instrument always involve errors. No measurement is free from errors.
Even if the precision of the equipment is adequate, no matter what its accuracy is, a discrepancy will always be observed between two measured results.
Since the accuracy of an instrument is measured in terms of its error, an understanding and evaluation of the errors is thus essential.

## Slide 18: STATIC ERROR
Static error is defined as the difference between the best measured value and the true value of the quantity. Then:
Es = Am - At
where, Es is static error OR absolute static error,
Am is measured value of quantity, and
At is true value of quantity.
The absolute value of error does not indicate precisely the accuracy of measurement.
For example, an error of 2A is negligible when the current being measured is of the order of 1000A; while the same error is highly significant if the current under measurement is of the order of 10A.

## Slide 19: STATIC ERROR
Thus another term relative static error is introduced.
The relative static error Er is the ratio of absolute static error to the true value of the quantity under measurement. Thus
    Er = Es/At
Percentage static error
% Er = Er x 100

## Slide 20: STATIC CORRECTION
Static Correction (C) is the difference between the true value and the measured value of the quantity
C = At – Am = -Es

## Slide 21: SCALE READABILITY
This indicates the closeness to with which the scale of an analog type of instrument can be read.
The readability of an instrument depends upon following factors:
i) Number of graduations
ii) Spacing between the graduations
iii) Size of the pointer
iv) Discriminating power of the observer
The readability is actually the number of significant figures in the instrument scale.
The higher the number of significant figures, the better would be the readability.

## Slide 22: REPEATABILITY AND REPRODUCIBILITY
Repeatability is the degree of closeness with which a given value may be repeatedly measured.
It is the closeness of output readings when the same input is applied repetitively over a short period of time.
The measurement is made on the same instrument, at the same location, by the same observer and under the same measurement conditions.
It may be specified in terms of units for a given period of time.

## Slide 23: REPEATABILITY AND REPRODUCIBILITY
Reproducibility relates to the closeness of output readings for the same input when there are changes in the method of measurement, observer, measuring instrument location, conditions of use and time of measurement.
Perfect reproducibility means that the instrument has no drift.
Drift means that with a given input the measured values vary with time.

## Slide 24: REPEATABILITY AND REPRODUCIBILITY
Reproducibility and Repeatability are a measure of closeness with which a given input may be measured over and over again.
The two terms cause confusion. Therefore, a distinction is made between the two terms.
Reproducibility is specified in terms of scale readings over a given period of time and location.
On the other hand, Repeatability is defined as the variation of scale reading and is random in nature.

## Slide 25: STATIC SENSITIVITY
The static sensitivity is the ratio of the magnitude of the output signal (qo) or response to the magnitude of input signal (qi) or the quantity being measured.
Measurement
System
INPUT (qi)
OUTPUT (qo)
[Force]
[Displacement]
Its units depend upon the type of input and output. If the output is in (mm) and the input is in (Newton) then its unit will be (mm/Newton).

## Slide 26: STATIC SENSITIVITY
Sometimes static sensitivity is also expressed as the ratio of the magnitude of the measured quantity (qi) to the magnitude of the response (qo).
Measurement
System
INPUT (qi)
OUTPUT (qo)
[Force]
[Displacement]
Thus the sensitivity expressed this way has the units of (Newton/mm). It is reciprocal of the sensitivity as defined above.
This ratio is defined as the inverse sensitivity or defection factor. Many manufacturers define the sensitivity of their instruments in terms of inverse sensitivity and still call it sensitivity.

## Slide 27: STATIC SENSITIVITY
When a calibration curve is linear [Fig (a) above], the slope of the calibration curve is constant.
The sensitivity is expressed as the slope of the calibration curve if the ordinates are expressed in actual units.
For this case the sensitivity is constant over the entire range of the instrument.

## Slide 28: STATIC SENSITIVITY
However, if the curve is not a straight line [Fig (b) above], the sensitivity varies with the input.
The sensitivity is expressed as the slope of the calibration curve if the ordinates are expressed in actual units.

## Slide 29: STATIC SENSITIVITY
In general, the static sensitivity at the operating point is defined as:
Similarly,
The sensitivity of an instrument should be high (greater scale readability & resolution) and therefore the instrument should not have a range greatly exceeding the value to be measured.
However, some margin should be kept for any accidental overloads.
Static sensitivity is determined by direct measurement (calibration) of (qo) against known input signal (qi).

## Slide 30: DRIFT
Drift is a departure in the output of the instrument over the period of time due to change in initial calibration conditions.
An instrument is said to have no drift if it produces same reading at different times for the same variation in the measured variable.
Drift is unrelated to the operating conditions or load.
The following factors could contribute towards the drift in the instruments:
    i) Wear and tear
    ii) Mechanical vibrations
    iii) Stresses developed in the parts of the instrument
    iv) Temperature variations
    v) Stray electric and magnetic fields
    vi) Thermal emf

## Slide 31: DRIFT
Drift can occur in the flow meters due to wear of nozzle or venturi.
It may occur in the resistance thermometer due to metal contamination etc.
Drift can occur due to change in spring stiffness with change in temperature.
Drift can be due to change in reading in spring balance because of backlash.
Drift is the reflection of change in sensitivity.

## Slide 32: DRIFT
Drift may be of any of the following types;
a) Zero drift: If the whole of instrument calibration shifts over by the same amount. It may be due to shifting of pointer or permanent set or backlash.
b) Span drift: If the calibration from zero upwards changes proportionately it is called span drift. It may be due to the change in spring gradient, change in Calibrated sensitivity and also known as sensitivity drift.
c) Zonal drift: When the drift occurs only over a portion of the span of the instrument it is called zonal drift.

## Slide 33: DRIFT
Drift is an undesirable quality in industrial instruments because it is rarely apparent and cannot be easily compensated for.
Thus, it must be carefully guarded against.
Effect of mechanical vibrations can be minimized by having proper mountings.
Temperature changes during the measurement process should be preferably avoided or otherwise be properly compensated for.

## Slide 34: LINEARITY
When the input-output points of the instrument are plotted, the resulting curve may not be linear.
This would be only if the output is proportional to input.
Linearity is the measure of maximum deviation of these points from the nominal straight line (may be obtained by least squares method).
The departure from the straight line relationship is non-linearity, but it is expressed as linearity of the instrument.
This departure from the straight line could be due to non-linear elements in the measuring system or the elastic after-effects of the mechanical system.

## Slide 35: LINEARITY
Linearity is expressed in many different ways:
    i) Independent Linearity:
    It is the maximum
    deviation from the
    straight line so placed
    as to minimize the
    maximum deviation .
    ii) Zero based linearity: It is the maximum deviation from the straight line joining the origin and so placed as to minimize the maximum deviation.
    iii)Terminal based linearity: It is the maximum deviation from the straight line joining both the end points of the curve.
Independent linearity

## Slide 36: LINEARITY

## Slide 37: LINEARITY
Linearity is considered to be one of the best characteristics of the measurement system, because of the convenience of scale reading.
A nonlinear relation does not lead to any inaccuracy, but it is better to choose the operating range instrument in such a way that the input-output relation is linear.
Lack of linearity thus does not necessarily degrade sensor performance.
If the nonlinearity can be modelled and an appropriate correction applied to the measurement before it is used for monitoring and control, the effect of the non-linearity can be eliminated.

## Slide 38: DEAD BAND AND DEAD TIME
Dead band, sometimes called a neutral zone, is an area of a signal range or band where no action occurs ( no output is obtainable), that is, the system is dead e.g. 10 g weight on a 10 kg balance.
It is the largest change in the input physical variable to which the measuring instrument does not respond.
In other words it is defined as the range of input values over which there is no change in output value.
It has also been referred to as Dead space or Dead zone.
In the analog instruments, it may occur due to friction in the instrument which does not allow pointer to move till sufficient force is developed to overcome the frictional loss.

## Slide 39: DEAD BAND AND DEAD TIME

PL
PH
x
Q
Q
x

Dead Band

## Slide 40: DEAD BAND AND DEAD TIME
Dead zone is specified by indicating it as the percentage of span range.
For example, if the calibration range of a thermometer is 100 to 300oC and the dead zone is specified as 0.1% of the span, then the change in temperature that must occur before it is detected by the instrument would be:
(300 -100) x 0.1 / 100 = 0.2oC
Dead band is different from hysteresis. Dead band is purposefully used in voltage regulators and other controllers to prevent oscillation or repeated activation-deactivation cycles.

## Slide 41: DEAD BAND AND DEAD TIME
Dead time is the time required by measuring instrument to begin to respond to a change in the measured variable.
It represents the time before the instrument begins to respond after the measured variable has changed.
The units of dead zone are the units of the variable, whereas, the units of the Dead time are the units of time.

## Slide 42: HYSTERESIS
Fig. A
Fig. B1
A
B1

## Slide 43: HYSTERESIS
B2
Fig. B2
Fig. B3
Fig. B4
B3
B4

## Slide 44: HYSTERESIS
Fig. B5
B2
B5
B5

## Slide 45: HYSTERESIS
Hysteresis is a phenomenon under which the measuring instrument shows different output effects during loading and unloading.
Hysteresis results from the inelastic quality of an element or device or for dry friction presence in the system.
It is the maximum difference between corresponding upscale and downscale outputs for the same input signal in any single test cycle, that is, the maximum difference in output at any given value of the measured variable within the specified range when approaching that point first with increasing the input and then decreasing the input.

## Slide 46: HYSTERESIS
Fig. B6: Input and Output Hysteresis

## Slide 47: HYSTERESIS
The numerical value of hysteresis is specified in terms of output or input values.
The sum of the two effects of dead band and the hysteresis is determined directly from the deviation values of a number of test cycles.
Hysteresis is determined by subtracting the value of dead band from the corresponding value of hysteresis plus dead band for a given input.
The maximum difference in the readings is reported as hysteresis. The difference may be expressed as a percentage of ideal output span.
For example, the hysteresis is reported as 0.2% of output span.
Hysteresis results due to the presence of several factors, such as, mechanical friction, motion in bearings, magnetic and thermal effects. This could also be due to a free play or looseness in the mechanisms.

## Slide 48: THRESHOLD
Threshold of a measuring instrument is the minimum value of input signal that is required to make a change or start from zero.
This is the minimum value below which no output change can be detected when the input is gradually increased from zero.
In digital system, the out put is displayed in incremental digits. Thus, in digital instruments the threshold is the minimum input signal which is necessary to produce at least one significant digit of output to indicate on the display.

## Slide 49: RESOLUTION
When an instrument is showing a particular output reading, there is a lower limit on the magnitude of the change in the input measured quantity that produces an observable change in the instrument output.
That means, when the input is slowly increased from some arbitrary input value, which is non-zero, the output does not change at all until certain increment is exceeded. This increment is called resolution or discrimination of the instrument. Thus, the resolution refers to the smallest change of input for which there will be a change output.

## Slide 50: RESOLUTION
In the analog instruments, the resolution is also determined by the observers ability to judge the position of pointer on the scale.
One of the major factors influencing the resolution of an instrument is how finely its output scale is divided into subdivisions. Using a car speedometer as an example again, this has subdivisions of typically 20 km/h. This means that when the needle is between the scale markings, we cannot estimate speed more accurately than to the nearest 5 km/h.

## Slide 51: THRESHOLD & RESOLUTION
The difference between threshold and the resolution of the measuring instrument could be understood this way.
Threshold defines the smallest measureable input, while the resolution defines the smallest measureable input change.
Both of these values may be expressed in terms of an actual value or as a fraction / percentage of the full scale value.

## Slide 52: NOISE
Noise is a random fluctuation in an electrical signal, a characteristic of all electronic circuits.
Noise generated by electronic devices varies greatly, as it can be produced by several different effects.
In general noise is an error or undesired random disturbance of a useful information signal, introduced before or after the detector and decoder.
The noise is a summation of unwanted or disturbing energy from natural and sometimes man-made sources.
The quality of a signal is often expressed quantitatively as the signal-to-noise ratio.

## Slide 53: NOISE
Noise may be defined as any signal that does not convey any useful information.
Extraneous disturbances generated in the measuring system itself or coming from outside, frequently constitute a background against which a signal may be read.
There are many sources of noise.
Noise may originate at the primary sensing device, in a communication channel or other intermediate links.
The noise may also be produced by indicating elements of the system.

## Slide 54: NOISE
The common sources of noise are given below:
i) Stray electrical and magnetic fields present in the neighbourhood of the instruments produce extraneous signals which tend to distort the original signal. The effects of these stray fields can be minimized by adequate shielding or relocation of the components of the instruments.
ii) Mechanical shocks and vibrations are another source of trouble. Their effect can be eliminated by proper mounting devices.

## Slide 55: You can contact us at:
Sanjib Kumar Acharyya
  9433229345 – Mobile & WhatsApp No
  sanjibacharyya24@gmail.com - Email
Saikat Mookherjee
  9051963640 – Mobile & WhatsApp No
  smookherjee13@gmail.com - Email